## Slide 1: Working with Self-Wicking Grids
Version 1.11

## Slide 2: Self-wicking grids produce thin ice without blotting
Bars coated with copper hydroxide nanowires
Removes the need for physically blotting away sample using filter paper
Can produce a thin film quickly, which can be beneficial for tackling air-water-interface denaturation/dissociation and preferential orientation problems
Glow discharge can be used to control the wicking process

## Slide 3: Grid preparation protocols
Sample + nanowire grid wicking behaviour is unknown at the start
First interaction requires wicking characterization to prepare grids with appropriate ice thickness
Use Two-stripe dispense mode to characterize wicking
When you know how your sample wicks on the grid, then use knowledge gained from characterization to move to One-stripe dispense mode

## Slide 4: Grid preparation protocols
Two-stripe (2-S) Dispense Mode
Wicking characterization stripe: To determine the appropriate wicking time for a particular grid and sample
Plunge stripe: To freeze the grid based on the wicking time determined using the wicking stripe
One-stripe (1-S) Dispense Mode
Plunge stripe only: To freeze the grid based on the parameters determined using the two-stripe dispense mode
Wicking characterization stripe
Plunge stripe

## Slide 5: Two-stripe Dispense Mode

## Slide 6: Step 1: Wicking Characterization (1st stripe)
Sample is sprayed on to the grid as it moves past the dispenser
Grid continues downward and arrives in front of the camera
A 2500ms video of the grid is captured in order to observe the wicking process
The grid is kept at room temp during characterization
Wicking characterization stripe
2-S Mode

## Slide 7: Step 1: Wicking Characterization (1st stripe)
An automated algorithm analyses the video and suggests a wicking time
44ms
300ms
100ms
2-S Mode

## Slide 8: Step 1: Wicking Characterization (1st stripe)
In this example, the algorithm identifies 100ms as the wicking time
44ms
300ms
100ms
2-S Mode

## Slide 9: Step 2: Plunge - Determine Plunge Offset for 2nd stripe
Wicking is usually slower when the grid is sprayed a second time
Going from characterization to plunge stripe, need to determine a “plunge offset”
Plunge Time = Wicking Time x Plunge offset
100ms
wicking time identified by algorithm
wicking characterization stripe
plunge stripe
2-S Mode
When the characterization stripe wicks quickly, use ~2x plunge offset as a starting point
  Ex. 100ms x 2 = 200ms
When the characterization stripe wicks more slowly, use a larger offset
  Ex. 500ms x 3 = 1500ms
Algorithm automatically suggests offset based on the wicking time from the characterization stripe
Plunge offset is not a well-defined variable, and will be different from one sample to another
2-S Mode

## Slide 10: Step 2: Plunge – 200ms dispense-to-plunge
The grid stays in front of the video camera until ~50ms before the selected plunge time
The video is used to decide to accept or reject the grid
It takes 44ms for the grid to travel from the dispense location into view of the camera
2-S Mode

## Slide 11: Getting the appropriate ice thickness
2 variables to adjust:
  Plunge offset
    Increasing the plunge offset allows wicking to happen for longer time
    More sample is wicked away with longer wicking time, resulting in grids with thinner ice
    Larger plunge offset/longer wicking time results in “slower” grids
  Glow discharge
    Increasing glow discharge increases the wicking speed
    Same degree of wicking occurs in less time with more glow discharge
increased glow discharge

## Slide 12: Getting the appropriate ice thickness
Ice thickness is inversely proportional to the degree of wicking
Degree of wicking is proportional to the wicking time
Underwicked
Reject
Underwicked
Reject
Underwicked
Acceptable
Good
Very Good
Overwicked
Acceptable
Time
Degree of Wicking
Ice thickness
Degree of Wicking
Wicking time
If the final frame of the wicking video before plunging appears as below, then how would we evaluate it?

## Slide 13: Case Study: achieving appropriately wicked grids
2 grids have been frozen using the two-stripe dispense mode
  Both are evaluated to be underwicked when a 2x plunge offset is applied
What should you do to get appropriately wicked grids?
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
2-S Mode

## Slide 14: Case Study: achieving appropriately wicked grids
More time for wicking, resulting in grids with thinner ice
Dispense-to-plunge time is longer
  may not be desirable for tackling sample denaturation/preferential orientation
Slower wicking grids (>300ms) generally require larger plunge offsets
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
3 | 12mA 40s | 450ms | 3x | 1350ms | Good
4 | 12mA 40s | 350ms | 3x | 1050ms | Good
2-S Mode
Option 1: Increase plunge offset

## Slide 15: Case Study: achieving appropriately wicked grids
  Wicking speed increases
  Dispense-to-plunge time decreases
  may be better for resolving AWI interface issues
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
3 | 12mA 80s | 100ms | 2x | 200ms | Good
4 | 12mA 80s | 80ms | 2x | 160ms | Good
2-S Mode
Option 2: Increase glow discharge

## Slide 16: Grid preparation protocols
Two-stripe (2-S) Dispense Mode
Wicking characterization stripe: To determine the appropriate wicking time and glow discharge parameters for a particular grid and sample
Plunge stripe: To freeze the grid based on the wicking time and glow discharge determined using the wicking stripe
One-stripe (1-S) Dispense Mode
Plunge stripe only: To freeze the grid based on the parameters determined using the two-stripe dispense mode
Wicking characterization stripe
Plunge stripe
2-S Mode

## Slide 17: One-stripe Dispense Mode

## Slide 18: Grid preparation protocols
Plunge stripe
Two-stripe (2-S) Dispense Mode
Wicking characterization stripe: To determine the appropriate wicking time and glow discharge parameters for a particular grid and sample
Plunge stripe: To freeze the grid based on the wicking time and glow discharge determined using the wicking stripe
One-stripe (1-S) Dispense Mode
Plunge stripe only: To freeze the grid based on the parameters determined using the two-stripe dispense mode
1-S Mode

## Slide 19: One-stripe Dispense Mode
No wicking characterization is performed. The grid is plunge frozen the first time the sample is dispensed onto the grid
A video of the wicking process is recorded by the camera on its way to the ethane cup
Accept or reject the grid based on the degree of wicking observed
1-S Mode

## Slide 20: One-stripe Dispense Mode
Dispense-to-plunge time can be nominated by user (54ms or 101ms – 2500ms)
Why use one-stripe dispense mode?
  Removes the need for a plunge offset
  Sample is sprayed onto the only grid once, achieving faster dispense-to-plunge time relative to two-stripe dispense mode
1-S Mode

## Slide 21: Case Study
  Two grids were plunged at 101ms, but both are underwicked.
  How can I get appropriately wicked grids using one-stripe dispense mode?
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | – | – | 101ms | Underwicked
2 | 12mA 40s | – | – | 101ms | Underwicked

## Slide 22: Case Study
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | – | – | 101ms | Underwicked
2 | 12mA 40s | – | – | 101ms | Underwicked
3 | 12mA 80s | – | – | 101ms | Overwicked
Option 1:
  Keep dispense-to-plunge time constant and increase glow discharge
    Wicking speed increases

## Slide 23: Case Study
  Consider the information we have now gained:
    Grid is wicking too slowly with 40s glow discharge (Grid #1, 2)
    Grid is wicking too quickly with 80s glow discharge (Grid #3)
  Reduce glow discharge time to 60s on the next grid (Grid #4)
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | – | – | 101ms | Underwicked
2 | 12mA 40s | – | – | 101ms | Underwicked
3 | 12mA 80s | – | – | 101ms | Overwicked
4 | 12mA 60s | – | – | 101ms | Good

## Slide 24: Case Study
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | – | – | 101ms | Underwicked
2 | 12mA 40s | – | – | 101ms | Underwicked
3 | 12mA 80s | – | – | 101ms | Overwicked
4 | 12mA 60s | – | – | 101ms | Good
5 | 12mA 80s | – | – | 54ms | Good
Option 2:
  Change the dispense-to-plunge time and hold glow discharge constant
    80s of glow discharge and 101ms dispense-to-plunge time resulted in an overwicked grid (#3)

## Slide 25: Freezing Strategy

## Slide 26: Two-stripe vs One-stripe Dispense Mode
 | Two-stripe | One-stripe
Advantages | Determine glow discharge required for the desired wicking speed while potentially getting useful grids | Faster dispense-to-plunge times Does not require a plunge offset
Disadvantages | Slower dispense-to-plunge times Variable offsets | Prior knowledge on glow discharge amount required

## Slide 27: Freezing Strategy
Use a mix of two-stripe and one-stripe dispense modes to optimize the glow discharge on-the-fly to get fast dispense-to-plunge times with appropriate ice thickness.
Use two-stripe dispense mode first
  Adjust glow discharge to achieve appropriate wicking times for the sample during wicking characterization step.
  Use offsets to still produce usable, but potentially slower grids
Switch to one-stripe dispense mode
  When the wicking characterization time is consistently in the target range switch to the one-stripe protocol to plunge in the target range
  Further increase glow discharge to go faster, e.g. from 101ms to 54ms

## Slide 28: Case Study
  How to do faster plunges to see if there is improvement in denaturation caused by the air-water-interface?
  When is the right time to switch to one-stripe protocol?
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
3 | 12mA 80s | 100ms | 2x | 200ms | Good
4 | 12mA 80s | 80ms | 2x | 160ms | Good

## Slide 29: Case Study
  Grid #3 and #4 with glow discharge of 12mA 80s both give wicking characterization times of 80-100ms
  Switch to one-stripe dispense mode on Grid #5 to plunge at 101ms, while keeping the same glow discharge settings
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
3 | 12mA 80s | 100ms | 2x | 200ms | Good
4 | 12mA 80s | 80ms | 2x | 160ms | Good
5 | 12mA 80s | – | – | 101ms | Good

## Slide 30: Case Study
  How to achieve the fastest dispense-to-plunge time possible (54ms)?
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
3 | 12mA 80s | 100ms | 2x | 200ms | Good
4 | 12mA 80s | 80ms | 2x | 160ms | Good
5 | 12mA 80s | – | – | 101ms | Good

## Slide 31: Case Study
  To plunge using 54ms one-stripe dispense mode, the grid needs to wick even faster than 101ms.
  Further increase glow discharge (i.e. from 80s to 100s) to speed up wicking
Grid | Glow Discharge | Wicking Characterization | Plunge Offset | Plunge | Evaluation
1 | 12mA 40s | 300ms | 2x | 600ms | Underwicked
2 | 12mA 40s | 400ms | 2x | 800ms | Underwicked
3 | 12mA 80s | 100ms | 2x | 200ms | Good
4 | 12mA 80s | 80ms | 2x | 160ms | Good
5 | 12mA 80s | – | – | 101ms | Good
6 | 12mA 100s | – | – | 54ms | Good

## Slide 32: Key Points Summary
Wicking can be assessed using videos and an automated algorithm
Film thickness (and resulting ice) can be controlled:
  Increasing glow discharge speeds up wicking, requiring less wicking time to achieve thin films
  Increasing plunge time gives more time for wicking, resulting in grids with thinner films
Use two-stripe dispense mode to determine the glow discharge required to achieve desired wicking speed while still producing grids
Use one-stripe dispense mode to achieve grids with short dispense-to-plunge times
Combine both two-stripe and one-stripe in the same freezing session to optimize the glow discharge needed to get appropriate ice thickness at short dispense-to-plunge time

## Slide 33: Quiz
Which of the following parameters can be adjusted to change the ice thickness achieved using self-wicking grids?
  Glow discharge
  Ethane temperature
  Plunge offset
  Dispense amplitude
When should I switch from two-stripe to one-stripe dispense mode?
  when grids are consistently underwicked
  when grids are consistently overwicked
  when grids are glow discharged for longer than 100s
  when consistently wicking time in the target range is achieved during wicking characterization
What are the benefits of one-stripe dispense mode?
  Can determine amount of glow discharge required for the desired wicking speed while potentially getting useful grids
  Does not require prior knowledge of glow discharge amount required to achieve the desired wicking speed
  Can freeze grids at short dispense-to-plunge time
  Does not require the use of offsets
What is the expected effect of increasing glow discharge applied?
  Increases wicking speed, grids have thinner ice
  Decreases wicking speed, grids have thicker ice
  Increases wicking time, grids have thinner ice
  Decreases wicking time, grids have thicker ice

## Slide 34: Quiz
Which of the following parameters can be adjusted to change the ice thickness achieved using self-wicking grids?
  Glow discharge
  Ethane temperature
  Plunge offset
  Dispense amplitude
When should I switch from two-stripe to one-stripe dispense mode?
  when grids are consistently underwicked
  when grids are consistently overwicked
  when grids are glow discharged for longer than 100s
  when consistently wicking time in the target range is achieved during wicking characterization
What are the benefits of one-stripe dispense mode?
  Can determine amount of glow discharge required for the desired wicking speed while potentially getting useful grids
  Does not require prior knowledge of glow discharge amount required to achieve the desired wicking speed
  Can freeze grids at short dispense-to-plunge time
  Does not require the use of offsets
What is the expected effect of increasing glow discharge applied?
  Increases wicking speed, grids have thinner ice
  Decreases wicking speed, grids have thicker ice
  Increases wicking time, grids have thinner ice
  Decreases wicking time, grids have thicker ice